## Per recording
recording | flow_type | curve_file | method | sample_count | duration_s | rmse_ml_s | mae_ml_s | r2 | correlation_r | mean_error_ml_s | maximum_absolute_error_ml_s | shape_rmse_normalized | reference_qmax_ml_s | predicted_qmax_ml_s | qmax_error_ml_s | absolute_qmax_error_ml_s | reference_volume_ml | predicted_volume_ml | volume_error_ml | absolute_volume_error_ml
Rec_data_83 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\001_Rec_data_83_curves.csv | Gradient Boosting | 151 | 19.65 | 3.5957 | 3.1369 | -0.027656 | 0.47758 | 0.92073 | 7.4598 | 0.31458 | 10.6287 | 12.0632 | 1.4345 | 1.4345 | 74.1055 | 92.2305 | 18.125 | 18.125
Rec_data_83 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\001_Rec_data_83_curves.csv | Calibration equation | 151 | 19.65 | 16.3138 | 13.6403 | -20.1534 | 0.40677 | 11.9658 | 31.7406 | 0.36059 | 10.6287 | 37.8849 | 27.2563 | 27.2563 | 74.1055 | 310.8299 | 236.7243 | 236.7243
Rec_data_83 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\001_Rec_data_83_curves.csv | Offset-corrected calibration | 151 | 19.65 | 4.2922 | 2.9561 | -0.46433 | 0.44935 | -1.8603 | 10.4716 | 0.38878 | 10.6287 | 14.353 | 3.7243 | 3.7243 | 74.1055 | 37.3355 | -36.77 | 36.77
Rec_data_84 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\002_Rec_data_84_curves.csv | Gradient Boosting | 140 | 18.4175 | 3.1824 | 2.6844 | 0.27036 | 0.87417 | 2.4052 | 7.8297 | 0.1548 | 11.9587 | 17.0862 | 5.1275 | 5.1275 | 85.5721 | 130.0762 | 44.5041 | 44.5041
Rec_data_84 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\002_Rec_data_84_curves.csv | Calibration equation | 140 | 18.4175 | 18.2279 | 15.3592 | -22.9364 | 0.82639 | 15.3345 | 33.8462 | 0.19649 | 11.9587 | 42.8668 | 30.9081 | 30.9081 | 85.5721 | 370.0694 | 284.4973 | 284.4973
Rec_data_84 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\002_Rec_data_84_curves.csv | Offset-corrected calibration | 140 | 18.4175 | 3.8568 | 2.781 | -0.071594 | 0.67512 | -0.84179 | 10.3142 | 0.30331 | 11.9587 | 19.3349 | 7.3761 | 7.3761 | 85.5721 | 69.9994 | -15.5727 | 15.5727
Rec_data_85 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\003_Rec_data_85_curves.csv | Gradient Boosting | 161 | 20.96 | 4.8609 | 4.2603 | -1.3214 | -0.16949 | 1.3281 | 8.2279 | 0.48337 | 9.4857 | 10.4244 | 0.93869 | 0.93869 | 59.5306 | 87.4485 | 27.9179 | 27.9179
Rec_data_85 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\003_Rec_data_85_curves.csv | Calibration equation | 161 | 20.96 | 16.6288 | 13.7686 | -26.1675 | -0.30869 | 10.325 | 31.3008 | 0.56258 | 9.4857 | 33.4914 | 24.0056 | 24.0056 | 59.5306 | 277.3183 | 217.7877 | 217.7877
Rec_data_85 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\003_Rec_data_85_curves.csv | Offset-corrected calibration | 161 | 20.96 | 4.562 | 3.075 | -1.0448 | -0.062399 | -1.563 | 9.4857 | 0.47487 | 9.4857 | 9.9594 | 0.47367 | 0.47367 | 59.5306 | 26.5875 | -32.9431 | 32.9431
Rec_data_86 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\004_Rec_data_86_curves.csv | Gradient Boosting | 143 | 18.602 | 3.4928 | 2.8054 | 0.30505 | 0.93704 | 2.7992 | 8.9623 | 0.12244 | 12.4373 | 19.1675 | 6.7302 | 6.7302 | 91.7513 | 144.0647 | 52.3134 | 52.3134
Rec_data_86 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\004_Rec_data_86_curves.csv | Calibration equation | 143 | 18.602 | 10.9014 | 8.7064 | -5.7698 | 0.90819 | 8.5455 | 22.1984 | 0.143 | 12.4373 | 32.4645 | 20.0271 | 20.0271 | 91.7513 | 251.8752 | 160.124 | 160.124
Rec_data_86 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\004_Rec_data_86_curves.csv | Offset-corrected calibration | 143 | 18.602 | 5.1204 | 3.8884 | -0.49356 | 0.61524 | -3.8884 | 12.2806 | 0.38781 | 12.4373 | 8.9325 | -3.5048 | 3.5048 | 91.7513 | 18.9507 | -72.8006 | 72.8006
Rec_data_87 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\005_Rec_data_87_curves.csv | Gradient Boosting | 131 | 17.16 | 3.0457 | 2.2803 | 0.48137 | 0.91753 | 2.1406 | 7.7467 | 0.1401 | 12.5532 | 18.9863 | 6.4331 | 6.4331 | 87.9356 | 124.7981 | 36.8625 | 36.8625
Rec_data_87 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\005_Rec_data_87_curves.csv | Calibration equation | 131 | 17.16 | 15.0235 | 11.6526 | -11.619 | 0.91733 | 11.4346 | 31.8687 | 0.13753 | 12.5532 | 44.0568 | 31.5036 | 31.5036 | 87.9356 | 285.6906 | 197.755 | 197.755
Rec_data_87 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\005_Rec_data_87_curves.csv | Offset-corrected calibration | 131 | 17.16 | 4.8076 | 3.8073 | -0.29224 | 0.70217 | -2.1151 | 10.3204 | 0.35893 | 12.5532 | 20.5248 | 7.9716 | 7.9716 | 87.9356 | 51.3886 | -36.547 | 36.547
Rec_data_88 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\006_Rec_data_88_curves.csv | Gradient Boosting | 137 | 17.816 | 2.3743 | 1.9246 | 0.87149 | 0.96134 | 1.3695 | 5.8225 | 0.091809 | 20.8906 | 25.3785 | 4.4879 | 4.4879 | 125.5198 | 149.9694 | 24.4496 | 24.4496
Rec_data_88 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\006_Rec_data_88_curves.csv | Calibration equation | 137 | 17.816 | 8.9108 | 6.6169 | -0.8101 | 0.95507 | 5.8834 | 19.5928 | 0.096072 | 20.8906 | 40.0619 | 19.1712 | 19.1712 | 125.5198 | 231.1805 | 105.6607 | 105.6607
Rec_data_88 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\006_Rec_data_88_curves.csv | Offset-corrected calibration | 137 | 17.816 | 6.5141 | 4.9993 | 0.032644 | 0.77943 | -4.9993 | 17.0505 | 0.2943 | 20.8906 | 16.5299 | -4.3607 | 4.3607 | 125.5198 | 35.87 | -89.6498 | 89.6498
Rec_data_89 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\007_Rec_data_89_curves.csv | Gradient Boosting | 139 | 17.94 | 1.7887 | 1.4307 | 0.92495 | 0.97519 | 1.0339 | 4.4994 | 0.078085 | 20.82 | 24.2206 | 3.4007 | 3.4007 | 138.924 | 157.4504 | 18.5264 | 18.5264
Rec_data_89 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\007_Rec_data_89_curves.csv | Calibration equation | 139 | 17.94 | 13.4191 | 10.3267 | -3.224 | 0.95609 | 9.3832 | 28.7991 | 0.096307 | 20.82 | 49.189 | 28.369 | 28.369 | 138.924 | 308.5239 | 169.5999 | 169.5999
Rec_data_89 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\007_Rec_data_89_curves.csv | Offset-corrected calibration | 139 | 17.94 | 5.9449 | 4.7165 | 0.17097 | 0.79973 | -4.0351 | 12.1325 | 0.29614 | 20.82 | 25.6571 | 4.8371 | 4.8371 | 138.924 | 66.0554 | -72.8686 | 72.8686
Rec_data_90 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\008_Rec_data_90_curves.csv | Gradient Boosting | 141 | 18.48 | 3.8753 | 3.0548 | 0.73625 | 0.98243 | 2.9921 | 7.5379 | 0.064992 | 21.9652 | 28.7902 | 6.825 | 6.825 | 172.012 | 227.5193 | 55.5073 | 55.5073
Rec_data_90 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\008_Rec_data_90_curves.csv | Calibration equation | 141 | 18.48 | 9.3866 | 7.3343 | -0.54737 | 0.98114 | 6.7877 | 17.1903 | 0.06875 | 21.9652 | 38.9659 | 17.0007 | 17.0007 | 172.012 | 298.4101 | 126.3981 | 126.3981
Rec_data_90 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\008_Rec_data_90_curves.csv | Offset-corrected calibration | 141 | 18.48 | 7.7317 | 6.3206 | -0.049867 | 0.82502 | -6.3206 | 15.7354 | 0.30428 | 21.9652 | 15.434 | -6.5312 | 6.5312 | 172.012 | 54.4383 | -117.5737 | 117.5737
Rec_data_91 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\009_Rec_data_91_curves.csv | Gradient Boosting | 140 | 18.348 | 1.5354 | 1.2422 | 0.95855 | 0.98194 | 0.5193 | 4.2842 | 0.06452 | 23.6963 | 23.7864 | 0.090125 | 0.090125 | 150.8965 | 160.3545 | 9.458 | 9.458
Rec_data_91 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\009_Rec_data_91_curves.csv | Calibration equation | 140 | 18.348 | 10.8013 | 8.0811 | -1.0515 | 0.97624 | 7.2866 | 23.2321 | 0.07278 | 23.6963 | 45.4857 | 21.7894 | 21.7894 | 150.8965 | 285.6053 | 134.7089 | 134.7089
Rec_data_91 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\009_Rec_data_91_curves.csv | Offset-corrected calibration | 140 | 18.348 | 6.1462 | 4.7478 | 0.33574 | 0.85648 | -4.7478 | 13.6132 | 0.25273 | 23.6963 | 21.9537 | -1.7425 | 1.7425 | 150.8965 | 63.21 | -87.6865 | 87.6865
Rec_data_92 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\010_Rec_data_92_curves.csv | Gradient Boosting | 152 | 19.932 | 2.2128 | 1.7542 | 0.91143 | 0.97908 | 1.4574 | 5.5418 | 0.072194 | 22.99 | 27.7496 | 4.7596 | 4.7596 | 154.5151 | 183.5173 | 29.0022 | 29.0022
Rec_data_92 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\010_Rec_data_92_curves.csv | Calibration equation | 152 | 19.932 | 8.9821 | 6.5229 | -0.45934 | 0.96909 | 5.3107 | 22.9206 | 0.093519 | 22.99 | 45.455 | 22.465 | 22.465 | 154.5151 | 260.7495 | 106.2344 | 106.2344
Rec_data_92 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\010_Rec_data_92_curves.csv | Offset-corrected calibration | 152 | 19.932 | 6.7623 | 5.1192 | 0.17284 | 0.80431 | -5.1192 | 16.1871 | 0.2902 | 22.99 | 21.923 | -1.0669 | 1.0669 | 154.5151 | 51.8599 | -102.6552 | 102.6552
Rec_data_93 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\011_Rec_data_93_curves.csv | Gradient Boosting | 137 | 17.952 | 3.4293 | 2.646 | 0.83657 | 0.9628 | 2.2969 | 10.7616 | 0.10593 | 25.0349 | 27.5418 | 2.5069 | 2.5069 | 155.0569 | 196.4084 | 41.3516 | 41.3516
Rec_data_93 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\011_Rec_data_93_curves.csv | Calibration equation | 137 | 17.952 | 7.6111 | 5.5683 | 0.19498 | 0.95402 | 4.8414 | 19.2463 | 0.11136 | 25.0349 | 36.8322 | 11.7974 | 11.7974 | 155.0569 | 242.6 | 87.5432 | 87.5432
Rec_data_93 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\011_Rec_data_93_curves.csv | Offset-corrected calibration | 137 | 17.952 | 8.1495 | 6.1917 | 0.077049 | 0.82938 | -6.1917 | 19.6285 | 0.25952 | 25.0349 | 13.3003 | -11.7346 | 11.7346 | 155.0569 | 43.1399 | -111.917 | 111.917
Rec_data_94 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\012_Rec_data_94_curves.csv | Gradient Boosting | 162 | 21.252 | 2.4683 | 2.0378 | 0.90971 | 0.96331 | 0.85184 | 6.8481 | 0.11274 | 25.879 | 23.3788 | -2.5001 | 2.5001 | 193.3104 | 211.4112 | 18.1007 | 18.1007
Rec_data_94 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\012_Rec_data_94_curves.csv | Calibration equation | 162 | 21.252 | 10.1857 | 8.2414 | -0.53747 | 0.92397 | 7.4481 | 21.8043 | 0.14931 | 25.879 | 39.5764 | 13.6974 | 13.6974 | 193.3104 | 352.6565 | 159.3461 | 159.3461
Rec_data_94 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\012_Rec_data_94_curves.csv | Offset-corrected calibration | 162 | 21.252 | 7.6729 | 6.12 | 0.12754 | 0.84912 | -6.12 | 18.11 | 0.24553 | 25.879 | 16.0445 | -9.8345 | 9.8345 | 193.3104 | 62.517 | -130.7935 | 130.7935
Rec_data_95 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\013_Rec_data_95_curves.csv | Gradient Boosting | 158 | 20.567 | 2.9817 | 2.3877 | 0.79915 | 0.9736 | 2.3279 | 6.715 | 0.081847 | 19.791 | 24.2101 | 4.4191 | 4.4191 | 161.6341 | 209.6549 | 48.0208 | 48.0208
Rec_data_95 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\013_Rec_data_95_curves.csv | Calibration equation | 158 | 20.567 | 13.5153 | 10.3995 | -3.1265 | 0.95322 | 9.9402 | 24.3219 | 0.11009 | 19.791 | 42.7379 | 22.9469 | 22.9469 | 161.6341 | 367.2708 | 205.6367 | 205.6367
Rec_data_95 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\013_Rec_data_95_curves.csv | Offset-corrected calibration | 158 | 20.567 | 4.488 | 3.51 | 0.54499 | 0.90545 | -3.4865 | 12.9044 | 0.2216 | 19.791 | 19.206 | -0.58502 | 0.58502 | 161.6341 | 89.5147 | -72.1194 | 72.1194
Rec_data_96 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\014_Rec_data_96_curves.csv | Gradient Boosting | 140 | 18.348 | 1.6824 | 1.2255 | 0.95978 | 0.98977 | -0.45312 | 4.2926 | 0.057845 | 26.1384 | 22.8634 | -3.275 | 3.275 | 185.7599 | 177.2668 | -8.4931 | 8.4931
Rec_data_96 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\014_Rec_data_96_curves.csv | Calibration equation | 140 | 18.348 | 13.5732 | 10.4935 | -1.6183 | 0.96972 | 9.7193 | 26.9899 | 0.086047 | 26.1384 | 50.2902 | 24.1518 | 24.1518 | 185.7599 | 365.4394 | 179.6795 | 179.6795
Rec_data_96 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\014_Rec_data_96_curves.csv | Offset-corrected calibration | 140 | 18.348 | 5.5063 | 4.4079 | 0.5691 | 0.91587 | -4.121 | 13.0202 | 0.21274 | 26.1384 | 26.7582 | 0.61987 | 0.61987 | 185.7599 | 109.6707 | -76.0892 | 76.0892
Rec_data_97 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\015_Rec_data_97_curves.csv | Gradient Boosting | 136 | 17.82 | 5.7565 | 4.8277 | 0.45623 | 0.99044 | 4.8277 | 10.7332 | 0.067952 | 23.3367 | 32.1164 | 8.7797 | 8.7797 | 153.8806 | 240.3375 | 86.4569 | 86.4569
Rec_data_97 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\015_Rec_data_97_curves.csv | Calibration equation | 136 | 17.82 | 20.1468 | 15.0575 | -5.6606 | 0.97929 | 14.0871 | 36.0872 | 0.089563 | 23.3367 | 57.7172 | 34.3804 | 34.3804 | 153.8806 | 406.7175 | 252.8369 | 252.8369
Rec_data_97 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\015_Rec_data_97_curves.csv | Offset-corrected calibration | 136 | 17.82 | 5.4311 | 3.8896 | 0.51596 | 0.96087 | 1.0134 | 12.5552 | 0.13343 | 23.3367 | 34.1852 | 10.8485 | 10.8485 | 153.8806 | 172.1239 | 18.2433 | 18.2433
Rec_data_98 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\016_Rec_data_98_curves.csv | Gradient Boosting | 143 | 18.602 | 1.4937 | 1.2318 | 0.96168 | 0.98363 | 0.57819 | 3.4793 | 0.067517 | 21.8083 | 21.9671 | 0.15879 | 0.15879 | 166.398 | 177.0885 | 10.6905 | 10.6905
Rec_data_98 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\016_Rec_data_98_curves.csv | Calibration equation | 143 | 18.602 | 14.0309 | 10.6454 | -2.381 | 0.97223 | 9.917 | 26.1543 | 0.086546 | 21.8083 | 46.7266 | 24.9182 | 24.9182 | 166.398 | 352.2201 | 185.8221 | 185.8221
Rec_data_98 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\016_Rec_data_98_curves.csv | Offset-corrected calibration | 143 | 18.602 | 4.936 | 3.5738 | 0.58158 | 0.90029 | -3.3423 | 13.9676 | 0.23142 | 21.8083 | 23.1946 | 1.3863 | 1.3863 | 166.398 | 103.8331 | -62.5649 | 62.5649
Rec_data_99 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\017_Rec_data_99_curves.csv | Gradient Boosting | 141 | 18.48 | 1.9628 | 1.4553 | 0.92136 | 0.97207 | 0.86183 | 5.2317 | 0.077011 | 21.6746 | 24.5797 | 2.9052 | 2.9052 | 147.4195 | 163.3189 | 15.8994 | 15.8994
Rec_data_99 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\017_Rec_data_99_curves.csv | Calibration equation | 141 | 18.48 | 7.2074 | 5.3534 | -0.060302 | 0.9395 | 2.1017 | 17.941 | 0.15388 | 21.6746 | 37.174 | 15.4994 | 15.4994 | 147.4195 | 186.6028 | 39.1833 | 39.1833
Rec_data_99 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\017_Rec_data_99_curves.csv | Offset-corrected calibration | 141 | 18.48 | 7.5957 | 6.0955 | -0.17762 | 0.80596 | -6.0955 | 17.744 | 0.30084 | 21.6746 | 13.642 | -8.0325 | 8.0325 | 147.4195 | 34.0368 | -113.3827 | 113.3827
Rec_data_100 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\018_Rec_data_100_curves.csv | Gradient Boosting | 133 | 17.292 | 3.2345 | 2.3661 | 0.87746 | 0.94233 | -0.16943 | 8.1014 | 0.12666 | 28.5587 | 23.7879 | -4.7708 | 4.7708 | 171.2317 | 168.1746 | -3.0571 | 3.0571
Rec_data_100 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\018_Rec_data_100_curves.csv | Calibration equation | 133 | 17.292 | 10.2675 | 7.3584 | -0.23473 | 0.92859 | 6.1537 | 27.1106 | 0.13764 | 28.5587 | 43.981 | 15.4224 | 15.4224 | 171.2317 | 278.5332 | 107.3015 | 107.3015
Rec_data_100 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\018_Rec_data_100_curves.csv | Offset-corrected calibration | 133 | 17.292 | 7.4915 | 5.8077 | 0.34268 | 0.84528 | -5.6037 | 18.1724 | 0.23484 | 28.5587 | 20.4491 | -8.1096 | 8.1096 | 171.2317 | 73.6836 | -97.5481 | 97.5481
Rec_data_101 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\019_Rec_data_101_curves.csv | Gradient Boosting | 156 | 20.46 | 4.005 | 3.5935 | -2.701 | -0.32885 | 1.8118 | 6.5289 | 0.50667 | 6.4613 | 8.4629 | 2.0016 | 2.0016 | 37.5624 | 74.7815 | 37.2191 | 37.2191
Rec_data_101 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\019_Rec_data_101_curves.csv | Calibration equation | 156 | 20.46 | 18.3758 | 14.7557 | -76.9137 | -0.46944 | 12.1694 | 34.2854 | 0.58521 | 6.4613 | 36.2418 | 29.7805 | 29.7805 | 37.5624 | 288.1776 | 250.6152 | 250.6152
Rec_data_101 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\019_Rec_data_101_curves.csv | Offset-corrected calibration | 156 | 20.46 | 3.533 | 2.4924 | -1.88 | -0.10776 | -0.43918 | 10.7534 | 0.43604 | 6.4613 | 12.7099 | 6.2486 | 6.2486 | 37.5624 | 28.5411 | -9.0212 | 9.0212
Rec_data_102 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\020_Rec_data_102_curves.csv | Gradient Boosting | 163 | 21.222 | 3.979 | 3.6676 | -1.0694 | 0.16049 | 1.9538 | 6.1161 | 0.42271 | 8.6753 | 9.4875 | 0.8122 | 0.8122 | 65.0804 | 106.7107 | 41.6303 | 41.6303
Rec_data_102 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\020_Rec_data_102_curves.csv | Calibration equation | 163 | 21.222 | 15.1756 | 13.265 | -29.1 | 0.0045764 | 11.7258 | 27.7122 | 0.46735 | 8.6753 | 29.7399 | 21.0646 | 21.0646 | 65.0804 | 315.4895 | 250.4092 | 250.4092
Rec_data_102 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\020_Rec_data_102_curves.csv | Offset-corrected calibration | 163 | 21.222 | 4.0719 | 3.0373 | -1.1671 | -0.052533 | -2.7596 | 8.6753 | 0.47593 | 8.6753 | 6.208 | -2.4674 | 2.4674 | 65.0804 | 6.1837 | -58.8967 | 58.8967
Rec_data_103 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\021_Rec_data_103_curves.csv | Gradient Boosting | 155 | 20.482 | 5.9648 | 5.1074 | -0.59794 | -0.13025 | 0.18656 | 12.5162 | 0.46198 | 14.4837 | 11.3092 | -3.1745 | 3.1745 | 88.5522 | 92.1999 | 3.6477 | 3.6477
Rec_data_103 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\021_Rec_data_103_curves.csv | Calibration equation | 155 | 20.482 | 19.3149 | 17.2638 | -15.7554 | -0.27342 | 13.6862 | 35.7936 | 0.53341 | 14.4837 | 37.9566 | 23.473 | 23.473 | 88.5522 | 369.4163 | 280.8641 | 280.8641
Rec_data_103 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\021_Rec_data_103_curves.csv | Offset-corrected calibration | 155 | 20.482 | 7.0473 | 4.9231 | -1.2306 | -0.06611 | -2.2807 | 14.4837 | 0.48722 | 14.4837 | 14.4247 | -0.058977 | 0.058977 | 88.5522 | 41.5845 | -46.9676 | 46.9676
Rec_data_104 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\022_Rec_data_104_curves.csv | Gradient Boosting | 145 | 18.864 | 1.4551 | 1.2132 | 0.95774 | 0.98452 | 0.75865 | 3.664 | 0.060091 | 23.1367 | 23.5632 | 0.42646 | 0.42646 | 150.0472 | 164.326 | 14.2788 | 14.2788
Rec_data_104 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\022_Rec_data_104_curves.csv | Calibration equation | 145 | 18.864 | 10.205 | 8.054 | -1.0787 | 0.9679 | 7.3053 | 21.2673 | 0.085557 | 23.1367 | 41.9931 | 18.8563 | 18.8563 | 150.0472 | 288.8292 | 138.782 | 138.782
Rec_data_104 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\022_Rec_data_104_curves.csv | Offset-corrected calibration | 145 | 18.864 | 6.2852 | 5.0945 | 0.21152 | 0.85673 | -5.0945 | 13.7668 | 0.24991 | 23.1367 | 18.4611 | -4.6756 | 4.6756 | 150.0472 | 53.3393 | -96.7078 | 96.7078
Rec_data_105 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\023_Rec_data_105_curves.csv | Gradient Boosting | 130 | 16.899 | 1.5897 | 1.228 | 0.94597 | 0.97524 | 0.35877 | 4.9101 | 0.074068 | 21.7882 | 21.316 | -0.47218 | 0.47218 | 124.6401 | 130.6251 | 5.985 | 5.985
Rec_data_105 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\023_Rec_data_105_curves.csv | Calibration equation | 130 | 16.899 | 8.9843 | 6.8087 | -0.72583 | 0.95387 | 6.0181 | 21.775 | 0.097715 | 21.7882 | 39.5722 | 17.784 | 17.784 | 124.6401 | 227.1735 | 102.5334 | 102.5334
Rec_data_105 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\023_Rec_data_105_curves.csv | Offset-corrected calibration | 130 | 16.899 | 6.7033 | 5.2538 | 0.039269 | 0.81182 | -5.2538 | 14.131 | 0.27715 | 21.7882 | 16.0403 | -5.748 | 5.748 | 124.6401 | 35.213 | -89.427 | 89.427
Rec_data_106 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\024_Rec_data_106_curves.csv | Gradient Boosting | 138 | 18.084 | 3.2415 | 2.4952 | 0.90333 | 0.96351 | 1.6325 | 9.9673 | 0.11093 | 32.4008 | 30.9093 | -1.4915 | 1.4915 | 204.1652 | 233.7394 | 29.5742 | 29.5742
Rec_data_106 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\024_Rec_data_106_curves.csv | Calibration equation | 138 | 18.084 | 18.0824 | 13.9105 | -2.0083 | 0.93077 | 13.0448 | 35.7759 | 0.14492 | 32.4008 | 59.969 | 27.5683 | 27.5683 | 204.1652 | 441.8771 | 237.712 | 237.712
Rec_data_106 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\024_Rec_data_106_curves.csv | Offset-corrected calibration | 138 | 18.084 | 5.246 | 3.8855 | 0.7468 | 0.92037 | -1.2267 | 12.3874 | 0.15378 | 32.4008 | 36.4371 | 4.0363 | 4.0363 | 204.1652 | 181.9087 | -22.2565 | 22.2565
Rec_data_107 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\025_Rec_data_107_curves.csv | Gradient Boosting | 136 | 17.82 | 2.1907 | 1.9107 | 0.66599 | 0.91389 | 1.4978 | 4.6389 | 0.1415 | 12.7248 | 13.8735 | 1.1487 | 1.1487 | 96.5411 | 123.3178 | 26.7767 | 26.7767
Rec_data_107 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\025_Rec_data_107_curves.csv | Calibration equation | 136 | 17.82 | 21.6665 | 18.0704 | -31.6712 | 0.87994 | 17.9895 | 36.8626 | 0.20176 | 12.7248 | 43.584 | 30.8593 | 30.8593 | 96.5411 | 419.5143 | 322.9732 | 322.9732
Rec_data_107 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\025_Rec_data_107_curves.csv | Offset-corrected calibration | 136 | 17.82 | 5.1718 | 4.0231 | -0.86157 | 0.797 | 1.3375 | 13.3307 | 0.24128 | 12.7248 | 20.0521 | 7.3273 | 7.3273 | 96.5411 | 120.5775 | 24.0364 | 24.0364
Rec_data_108 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\026_Rec_data_108_curves.csv | Gradient Boosting | 144 | 18.733 | 5.1459 | 4.554 | -0.55033 | 0.29566 | 1.5628 | 9.7091 | 0.36952 | 12.0478 | 15.3144 | 3.2665 | 3.2665 | 77.2581 | 106.6386 | 29.3805 | 29.3805
Rec_data_108 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\026_Rec_data_108_curves.csv | Calibration equation | 144 | 18.733 | 15.6204 | 12.6612 | -13.2851 | 0.25172 | 10.3969 | 32.8397 | 0.4045 | 12.0478 | 38.1818 | 26.134 | 26.134 | 77.2581 | 273.4295 | 196.1715 | 196.1715
Rec_data_108 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\026_Rec_data_108_curves.csv | Offset-corrected calibration | 144 | 18.733 | 5.6545 | 3.809 | -0.87188 | 0.1606 | -2.5738 | 12.0478 | 0.4551 | 12.0478 | 14.6498 | 2.602 | 2.602 | 77.2581 | 28.75 | -48.5081 | 48.5081
Rec_data_109 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\027_Rec_data_109_curves.csv | Gradient Boosting | 149 | 19.388 | 3.6168 | 2.8741 | 0.75541 | 0.97439 | 2.8741 | 8.7585 | 0.087079 | 21.4409 | 29.0622 | 7.6214 | 7.6214 | 186.5485 | 242.4699 | 55.9214 | 55.9214
Rec_data_109 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\027_Rec_data_109_curves.csv | Calibration equation | 149 | 19.388 | 6.3033 | 4.9489 | 0.25709 | 0.96407 | 3.8885 | 14.4052 | 0.10974 | 21.4409 | 35.0185 | 13.5776 | 13.5776 | 186.5485 | 262.5133 | 75.9648 | 75.9648
Rec_data_109 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\027_Rec_data_109_curves.csv | Offset-corrected calibration | 149 | 19.388 | 10.1481 | 8.299 | -0.92561 | 0.66835 | -8.299 | 18.4356 | 0.42127 | 21.4409 | 11.4866 | -9.9543 | 9.9543 | 186.5485 | 24.625 | -161.9235 | 161.9235
Rec_data_110 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\028_Rec_data_110_curves.csv | Gradient Boosting | 138 | 17.947 | 1.3038 | 1.0887 | 0.92287 | 0.98053 | 0.90948 | 3.0964 | 0.06119 | 16.3471 | 17.7227 | 1.3757 | 1.3757 | 120.4096 | 136.7416 | 16.332 | 16.332
Rec_data_110 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\028_Rec_data_110_curves.csv | Calibration equation | 138 | 17.947 | 13.1717 | 11.1675 | -6.8728 | 0.97061 | 11.0866 | 22.5478 | 0.088895 | 16.3471 | 38.6342 | 22.2871 | 22.2871 | 120.4096 | 320.8776 | 200.468 | 200.468
Rec_data_110 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\028_Rec_data_110_curves.csv | Offset-corrected calibration | 138 | 17.947 | 5.4338 | 4.4652 | -0.33985 | 0.7694 | -4.4652 | 10.1219 | 0.32732 | 16.3471 | 15.1022 | -1.2449 | 1.2449 | 120.4096 | 39.7319 | -80.6777 | 80.6777
Rec_data_111 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\029_Rec_data_111_curves.csv | Gradient Boosting | 141 | 18.48 | 2.2794 | 1.77 | 0.91207 | 0.96526 | -0.22533 | 5.4508 | 0.099428 | 24.7912 | 20.8593 | -3.9319 | 3.9319 | 166.768 | 162.4535 | -4.3145 | 4.3145
Rec_data_111 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\029_Rec_data_111_curves.csv | Calibration equation | 141 | 18.48 | 6.1129 | 4.7952 | 0.36754 | 0.9468 | 4.0633 | 13.1274 | 0.12364 | 24.7912 | 31.8488 | 7.0576 | 7.0576 | 166.768 | 242.4478 | 75.6798 | 75.6798
Rec_data_111 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\029_Rec_data_111_curves.csv | Offset-corrected calibration | 141 | 18.48 | 9.6057 | 7.688 | -0.56165 | 0.82591 | -7.688 | 19.4859 | 0.27666 | 24.7912 | 8.3168 | -16.4744 | 16.4744 | 166.768 | 23.733 | -143.035 | 143.035
Rec_data_112 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\030_Rec_data_112_curves.csv | Gradient Boosting | 127 | 16.506 | 2.8728 | 2.3157 | 0.63605 | 0.96988 | 2.3157 | 7.4725 | 0.079852 | 15.5969 | 21.0411 | 5.4442 | 5.4442 | 87.0901 | 125.4713 | 38.3812 | 38.3812
Rec_data_112 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\030_Rec_data_112_curves.csv | Calibration equation | 127 | 16.506 | 14.5104 | 10.9735 | -8.285 | 0.95686 | 10.5988 | 31.2713 | 0.093468 | 15.5969 | 45.2174 | 29.6205 | 29.6205 | 87.0901 | 263.4602 | 176.3701 | 176.3701
Rec_data_112 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\030_Rec_data_112_curves.csv | Offset-corrected calibration | 127 | 16.506 | 3.9709 | 3.163 | 0.30466 | 0.83647 | -2.112 | 7.7393 | 0.25576 | 15.5969 | 21.6854 | 6.0885 | 6.0885 | 87.0901 | 51.9912 | -35.0989 | 35.0989
Rec_data_113 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\031_Rec_data_113_curves.csv | Gradient Boosting | 136 | 17.82 | 2.8098 | 2.1624 | 0.9172 | 0.97143 | 1.5862 | 7.4717 | 0.11173 | 27.6872 | 26.765 | -0.92227 | 0.92227 | 189.9486 | 218.2731 | 28.3244 | 28.3244
Rec_data_113 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\031_Rec_data_113_curves.csv | Calibration equation | 136 | 17.82 | 11.1202 | 8.4346 | -0.29689 | 0.96324 | 7.3396 | 20.9123 | 0.11061 | 27.6872 | 44.2299 | 16.5426 | 16.5426 | 189.9486 | 321.7827 | 131.8341 | 131.8341
Rec_data_113 | Normal | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\031_Rec_data_113_curves.csv | Offset-corrected calibration | 136 | 17.82 | 6.5365 | 5.1636 | 0.55191 | 0.92257 | -5.1636 | 15.8229 | 0.18689 | 27.6872 | 20.6979 | -6.9893 | 6.9893 | 189.9486 | 97.3241 | -92.6245 | 92.6245
Interrupted_15 | Interrupted | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\032_Interrupted_15_curves.csv | Gradient Boosting | 184 | 24.156 | 3.2776 | 2.5537 | -2.4646 | 0.88303 | 2.5537 | 9.0607 | 0.14354 | 5.8565 | 12.6386 | 6.7822 | 6.7822 | 46.7702 | 108.7133 | 61.9431 | 61.9431
Interrupted_15 | Interrupted | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\032_Interrupted_15_curves.csv | Calibration equation | 184 | 24.156 | 9.6001 | 6.9092 | -28.7243 | 0.90141 | 6.8323 | 22.4301 | 0.13905 | 5.8565 | 25.9686 | 20.1121 | 20.1121 | 46.7702 | 212.6545 | 165.8843 | 165.8843
Interrupted_15 | Interrupted | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\032_Interrupted_15_curves.csv | Offset-corrected calibration | 184 | 24.156 | 2.5598 | 1.8773 | -1.1133 | 0.15441 | -1.8773 | 5.8565 | 0.43603 | 5.8565 | 2.4366 | -3.4198 | 3.4198 | 46.7702 | 1.1861 | -45.584 | 45.584
Interrupted_16 | Interrupted | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\033_Interrupted_16_curves.csv | Gradient Boosting | 188 | 24.684 | 4.1388 | 3.3468 | -1.0622 | 0.78121 | 3.1145 | 8.9628 | 0.23898 | 8.0142 | 13.0497 | 5.0355 | 5.0355 | 88.9767 | 166.1961 | 77.2194 | 77.2194
Interrupted_16 | Interrupted | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\033_Interrupted_16_curves.csv | Calibration equation | 188 | 24.684 | 18.0839 | 14.9124 | -38.3684 | 0.76169 | 14.7623 | 29.5862 | 0.26154 | 8.0142 | 34.6561 | 26.6419 | 26.6419 | 88.9767 | 455.3403 | 366.3636 | 366.3636
Interrupted_16 | Interrupted | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\033_Interrupted_16_curves.csv | Offset-corrected calibration | 188 | 24.684 | 2.8965 | 2.1382 | -0.01 | 0.67189 | -0.42886 | 7.4682 | 0.32952 | 8.0142 | 11.1241 | 3.11 | 3.11 | 88.9767 | 78.3555 | -10.6212 | 10.6212
Interrupted_17 | Interrupted | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\034_Interrupted_17_curves.csv | Gradient Boosting | 200 | 26.268 | 2.4205 | 1.9067 | 0.20767 | 0.66087 | 0.84137 | 6.1292 | 0.27731 | 8.4282 | 8.7269 | 0.29872 | 0.29872 | 85.6405 | 107.8405 | 22.2 | 22.2
Interrupted_17 | Interrupted | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\034_Interrupted_17_curves.csv | Calibration equation | 200 | 26.268 | 8.4806 | 6.9475 | -8.726 | 0.63124 | 6.3555 | 16.1366 | 0.30155 | 8.4282 | 20.1034 | 11.6752 | 11.6752 | 85.6405 | 253.476 | 167.8355 | 167.8355
Interrupted_17 | Interrupted | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\034_Interrupted_17_curves.csv | Offset-corrected calibration | 200 | 26.268 | 4.2344 | 3.2458 | -1.4247 |  | -3.2458 | 8.4282 |  | 8.4282 | 0 | -8.4282 | 8.4282 | 85.6405 | 0 | -85.6405 | 85.6405
Interrupted_18 | Interrupted | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\035_Interrupted_18_curves.csv | Gradient Boosting | 193 | 25.344 | 2.386 | 1.923 | -0.27217 | 0.79814 | 1.6891 | 4.8467 | 0.2116 | 6.6668 | 9.1434 | 2.4766 | 2.4766 | 69.1794 | 112.1863 | 43.0069 | 43.0069
Interrupted_18 | Interrupted | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\035_Interrupted_18_curves.csv | Calibration equation | 193 | 25.344 | 19.9795 | 16.7167 | -88.2052 | 0.72999 | 16.6579 | 32.706 | 0.26001 | 6.6668 | 37.8358 | 31.169 | 31.169 | 69.1794 | 493.5968 | 424.4174 | 424.4174
Interrupted_18 | Interrupted | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\035_Interrupted_18_curves.csv | Offset-corrected calibration | 193 | 25.344 | 3.7065 | 2.6465 | -2.0701 | 0.66137 | 0.70592 | 9.1741 | 0.31402 | 6.6668 | 14.3038 | 7.637 | 7.637 | 69.1794 | 87.2043 | 18.0249 | 18.0249
Interrupted_19 | Interrupted | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\036_Interrupted_19_curves.csv | Gradient Boosting | 173 | 22.704 | 2.2197 | 1.5835 | 0.75485 | 0.88666 | 0.59072 | 5.5741 | 0.17259 | 12.4659 | 12.8822 | 0.41628 | 0.41628 | 100.8944 | 114.3495 | 13.4552 | 13.4552
Interrupted_19 | Interrupted | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\036_Interrupted_19_curves.csv | Calibration equation | 173 | 22.704 | 6.8789 | 4.629 | -1.3545 | 0.90068 | 4.0242 | 16.1078 | 0.16252 | 12.4659 | 25.7564 | 13.2905 | 13.2905 | 100.8944 | 192.8495 | 91.9551 | 91.9551
Interrupted_19 | Interrupted | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\036_Interrupted_19_curves.csv | Offset-corrected calibration | 173 | 22.704 | 6.2057 | 4.3667 | -0.91623 | 0.26528 | -4.3667 | 12.4659 | 0.48123 | 12.4659 | 2.2244 | -10.2415 | 10.2415 | 100.8944 | 1.2343 | -99.66 | 99.66
Interrupted_20 | Interrupted | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\037_Interrupted_20_curves.csv | Gradient Boosting | 201 | 26.4 | 3.1965 | 2.7421 | -8.734 | 0.89226 | 2.7421 | 5.6526 | 0.15105 | 3.1659 | 8.6617 | 5.4957 | 5.4957 | 36.8268 | 109.527 | 72.7002 | 72.7002
Interrupted_20 | Interrupted | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\037_Interrupted_20_curves.csv | Calibration equation | 201 | 26.4 | 12.0031 | 9.6733 | -136.2521 | 0.86819 | 9.6302 | 22.969 | 0.1637 | 3.1659 | 26.0082 | 22.8423 | 22.8423 | 36.8268 | 292.3539 | 255.5271 | 255.5271
Interrupted_20 | Interrupted | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\037_Interrupted_20_curves.csv | Offset-corrected calibration | 201 | 26.4 | 1.677 | 1.3488 | -1.6793 | 0.23459 | -1.3488 | 3.1271 | 0.52799 | 3.1659 | 2.4763 | -0.68966 | 0.68966 | 36.8268 | 1.0589 | -35.7679 | 35.7679
Slow_18 | Slow | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\038_Slow_18_curves.csv | Gradient Boosting | 221 | 29.04 | 1.3323 | 1.0037 | -0.11048 | 0.57096 | 0.55029 | 3.39 | 0.33742 | 3.6258 | 6.3629 | 2.737 | 2.737 | 65.0532 | 81.0772 | 16.0239 | 16.0239
Slow_18 | Slow | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\038_Slow_18_curves.csv | Calibration equation | 221 | 29.04 | 10.5438 | 8.4385 | -68.5557 | 0.38855 | 7.6055 | 28.2918 | 0.45396 | 3.6258 | 31.3159 | 27.6901 | 27.6901 | 65.0532 | 286.9382 | 221.885 | 221.885
Slow_18 | Slow | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\038_Slow_18_curves.csv | Offset-corrected calibration | 221 | 29.04 | 2.5663 | 2.2181 | -3.1205 | 0.10296 | -2.069 | 4.7599 | 0.6941 | 3.6258 | 7.7839 | 4.1581 | 4.1581 | 65.0532 | 4.7141 | -60.3392 | 60.3392
Slow_19 | Slow | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\039_Slow_19_curves.csv | Gradient Boosting | 197 | 25.872 | 7.2982 | 6.044 | -7.5714 | 0.060193 | 5.7222 | 13.0718 | 0.49573 | 6.3935 | 13.6635 | 7.27 | 7.27 | 82.1385 | 230.8518 | 148.7133 | 148.7133
Slow_19 | Slow | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\039_Slow_19_curves.csv | Calibration equation | 197 | 25.872 | 9.1578 | 7.9528 | -12.4961 | 0.091645 | 7.3558 | 15.8333 | 0.49525 | 6.3935 | 16.51 | 10.1165 | 10.1165 | 82.1385 | 273.4392 | 191.3006 | 191.3006
Slow_19 | Slow | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\039_Slow_19_curves.csv | Offset-corrected calibration | 197 | 25.872 | 4.0245 | 3.1595 | -1.6064 |  | -3.1595 | 6.3935 |  | 6.3935 | 0 | -6.3935 | 6.3935 | 82.1385 | 0 | -82.1385 | 82.1385
Slow_20 | Slow | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\040_Slow_20_curves.csv | Gradient Boosting | 198 | 26.004 | 3.6562 | 3.0968 | -8.4729 | 0.67799 | 3.0968 | 6.906 | 0.25184 | 3.5523 | 9.6684 | 6.1161 | 6.1161 | 53.0327 | 133.9073 | 80.8746 | 80.8746
Slow_20 | Slow | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\040_Slow_20_curves.csv | Calibration equation | 198 | 26.004 | 7.8373 | 6.4827 | -42.5265 | 0.69066 | 6.3223 | 14.4921 | 0.26878 | 3.5523 | 17.083 | 13.5307 | 13.5307 | 53.0327 | 218.2805 | 165.2478 | 165.2478
Slow_20 | Slow | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\040_Slow_20_curves.csv | Offset-corrected calibration | 198 | 26.004 | 2.3515 | 2.0294 | -2.9184 |  | -2.0294 | 3.5523 |  | 3.5523 | 0 | -3.5523 | 3.5523 | 53.0327 | 0 | -53.0327 | 53.0327
Slow_21 | Slow | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\041_Slow_21_curves.csv | Gradient Boosting | 231 | 30.36 | 1.7686 | 1.4137 | -0.65508 | 0.84632 | 1.3726 | 4.0029 | 0.18375 | 4.0491 | 6.4187 | 2.3696 | 2.3696 | 64.562 | 106.3783 | 41.8163 | 41.8163
Slow_21 | Slow | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\041_Slow_21_curves.csv | Calibration equation | 231 | 30.36 | 13.4193 | 10.3122 | -94.2825 | 0.83243 | 10.0982 | 25.8589 | 0.22177 | 4.0491 | 28.8516 | 24.8025 | 24.8025 | 64.562 | 372.4837 | 307.9217 | 307.9217
Slow_21 | Slow | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\041_Slow_21_curves.csv | Offset-corrected calibration | 231 | 30.36 | 2.3465 | 1.9174 | -1.9133 | 0.22758 | -1.7993 | 4.0488 | 0.57982 | 4.0491 | 5.3196 | 1.2705 | 1.2705 | 64.562 | 9.7048 | -54.8572 | 54.8572
Slow_22 | Slow | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\042_Slow_22_curves.csv | Gradient Boosting | 273 | 35.904 | 7.282 | 6.7473 | -49.5228 | 0.20468 | 6.7473 | 12.3508 | 0.31776 | 3.5783 | 14.4581 | 10.8799 | 10.8799 | 92.0303 | 335.0408 | 243.0105 | 243.0105
Slow_22 | Slow | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\042_Slow_22_curves.csv | Calibration equation | 273 | 35.904 | 10.2084 | 9.3348 | -98.2899 | 0.2421 | 9.2043 | 17.3884 | 0.34552 | 3.5783 | 20.3356 | 16.7573 | 16.7573 | 92.0303 | 423.7251 | 331.6948 | 331.6948
Slow_22 | Slow | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\042_Slow_22_curves.csv | Offset-corrected calibration | 273 | 35.904 | 2.7519 | 2.5541 | -6.2151 |  | -2.5541 | 3.5783 |  | 3.5783 | 0 | -3.5783 | 3.5783 | 92.0303 | 0 | -92.0303 | 92.0303
Slow_23 | Slow | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\043_Slow_23_curves.csv | Gradient Boosting | 206 | 27.06 | 3.8438 | 3.5686 | -3.6046 | 0.91003 | 3.5686 | 5.773 | 0.15563 | 4.9395 | 9.7801 | 4.8406 | 4.8406 | 95.9068 | 192.8999 | 96.9931 | 96.9931
Slow_23 | Slow | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\043_Slow_23_curves.csv | Calibration equation | 206 | 27.06 | 20.423 | 19.0318 | -128.9903 | 0.90553 | 19.0317 | 26.4911 | 0.15961 | 4.9395 | 31.4307 | 26.4911 | 26.4911 | 95.9068 | 613.3681 | 517.4613 | 517.4613
Slow_23 | Slow | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\043_Slow_23_curves.csv | Offset-corrected calibration | 206 | 27.06 | 1.8019 | 1.3253 | -0.011893 | 0.71244 | -0.80757 | 4.9123 | 0.44997 | 4.9395 | 7.8987 | 2.9592 | 2.9592 | 95.9068 | 73.9716 | -21.9352 | 21.9352
Slow_24 | Slow | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\044_Slow_24_curves.csv | Gradient Boosting | 204 | 26.796 | 3.7661 | 3.3634 | -1.7181 | 0.88348 | 3.3634 | 7.3944 | 0.17219 | 6.3126 | 11.4835 | 5.1709 | 5.1709 | 112.8331 | 203.3429 | 90.5099 | 90.5099
Slow_24 | Slow | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\044_Slow_24_curves.csv | Calibration equation | 204 | 26.796 | 31.2652 | 28.7274 | -186.3314 | 0.87512 | 28.7274 | 42.7862 | 0.17706 | 6.3126 | 48.7292 | 42.4166 | 42.4166 | 112.8331 | 886.3991 | 773.5661 | 773.5661
Slow_24 | Slow | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\044_Slow_24_curves.csv | Offset-corrected calibration | 204 | 26.796 | 10.3969 | 8.8539 | -19.7157 | 0.85424 | 8.5934 | 19.2543 | 0.24477 | 6.3126 | 25.1972 | 18.8846 | 18.8846 | 112.8331 | 344.2593 | 231.4262 | 231.4262
Slow_25 | Slow | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\045_Slow_25_curves.csv | Gradient Boosting | 166 | 21.78 | 2.9455 | 2.8297 | -11.8703 | 0.63054 | 2.8297 | 4.8075 | 0.16982 | 3.8235 | 7.8806 | 4.0571 | 4.0571 | 63.6581 | 125.3713 | 61.7132 | 61.7132
Slow_25 | Slow | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\045_Slow_25_curves.csv | Calibration equation | 166 | 21.78 | 14.451 | 14.1428 | -308.7856 | 0.60801 | 14.1391 | 20.1445 | 0.17776 | 3.8235 | 23.968 | 20.1445 | 20.1445 | 63.6581 | 372.7127 | 309.0546 | 309.0546
Slow_25 | Slow | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\045_Slow_25_curves.csv | Offset-corrected calibration | 166 | 21.78 | 3.015 | 2.9021 | -12.485 | 0.11784 | -2.9021 | 3.8165 | 0.78296 | 3.8235 | 0.43608 | -3.3874 | 3.3874 | 63.6581 | 0.089539 | -63.5685 | 63.5685
Staccato_18 | Staccato | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\046_Staccato_18_curves.csv | Gradient Boosting | 168 | 22.044 | 1.4165 | 1.0537 | 0.80547 | 0.92649 | 0.72113 | 3.6216 | 0.14311 | 9.7521 | 9.8174 | 0.065329 | 0.065329 | 108.1917 | 124.1618 | 15.9701 | 15.9701
Staccato_18 | Staccato | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\046_Staccato_18_curves.csv | Calibration equation | 168 | 22.044 | 20.3203 | 17.8604 | -39.0327 | 0.9091 | 17.8589 | 29.2891 | 0.17061 | 9.7521 | 38.0573 | 28.3052 | 28.3052 | 108.1917 | 504.2397 | 396.048 | 396.048
Staccato_18 | Staccato | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\046_Staccato_18_curves.csv | Offset-corrected calibration | 168 | 22.044 | 2.896 | 2.3323 | 0.18688 | 0.83853 | -0.13755 | 6.3131 | 0.25895 | 9.7521 | 14.5253 | 4.7732 | 4.7732 | 108.1917 | 105.1894 | -3.0023 | 3.0023
Staccato_19 | Staccato | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\047_Staccato_19_curves.csv | Gradient Boosting | 183 | 24.024 | 6.6407 | 5.9018 | -4.2832 | 0.94423 | 5.9018 | 10.6319 | 0.11405 | 8.9627 | 18.6669 | 9.7043 | 9.7043 | 114.3081 | 256.7647 | 142.4566 | 142.4566
Staccato_19 | Staccato | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\047_Staccato_19_curves.csv | Calibration equation | 183 | 24.024 | 26.4234 | 23.1819 | -82.6468 | 0.92427 | 23.1657 | 38.5915 | 0.1459 | 8.9627 | 46.5108 | 37.5481 | 37.5481 | 114.3081 | 673.8573 | 559.5492 | 559.5492
Staccato_19 | Staccato | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\047_Staccato_19_curves.csv | Offset-corrected calibration | 183 | 24.024 | 7.1848 | 5.6267 | -5.1844 | 0.89134 | 4.6084 | 15.0596 | 0.1987 | 8.9627 | 22.9788 | 14.0162 | 14.0162 | 114.3081 | 225.6709 | 111.3628 | 111.3628
Staccato_20 | Staccato | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\048_Staccato_20_curves.csv | Gradient Boosting | 171 | 22.44 | 2.3281 | 2.0351 | 0.26272 | 0.97869 | 2.0351 | 4.0538 | 0.069877 | 8.2267 | 11.669 | 3.4423 | 3.4423 | 101.155 | 147.0291 | 45.8741 | 45.8741
Staccato_20 | Staccato | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\048_Staccato_20_curves.csv | Calibration equation | 171 | 22.44 | 26.5493 | 23.0756 | -94.8788 | 0.95614 | 22.8906 | 41.118 | 0.098775 | 8.2267 | 48.789 | 40.5623 | 40.5623 | 101.155 | 617.8639 | 516.7089 | 516.7089
Staccato_20 | Staccato | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\048_Staccato_20_curves.csv | Offset-corrected calibration | 171 | 22.44 | 7.5619 | 5.6497 | -6.7782 | 0.93938 | 4.6767 | 17.586 | 0.21572 | 8.2267 | 25.2571 | 17.0304 | 17.0304 | 101.155 | 206.7406 | 105.5856 | 105.5856
Staccato_21 | Staccato | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\049_Staccato_21_curves.csv | Gradient Boosting | 172 | 22.572 | 2.7507 | 1.9378 | 0.54382 | 0.86975 | 1.5262 | 7.6373 | 0.16138 | 12.6133 | 15.388 | 2.7747 | 2.7747 | 133.1485 | 167.7376 | 34.5891 | 34.5891
Staccato_21 | Staccato | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\049_Staccato_21_curves.csv | Calibration equation | 172 | 22.572 | 18.328 | 15.5014 | -19.2527 | 0.87188 | 15.3233 | 31.7406 | 0.1733 | 12.6133 | 40.8607 | 28.2473 | 28.2473 | 133.1485 | 480.9473 | 347.7988 | 347.7988
Staccato_21 | Staccato | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\049_Staccato_21_curves.csv | Offset-corrected calibration | 172 | 22.572 | 3.8667 | 3.0952 | 0.09857 | 0.74649 | -1.4825 | 8.2086 | 0.30962 | 12.6133 | 17.3287 | 4.7154 | 4.7154 | 133.1485 | 99.5369 | -33.6116 | 33.6116
Staccato_22 | Staccato | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\050_Staccato_22_curves.csv | Gradient Boosting | 188 | 24.497 | 2.8039 | 2.1604 | 0.41919 | 0.92835 | 2.1537 | 6.5453 | 0.11874 | 12.2686 | 15.6049 | 3.3362 | 3.3362 | 141.2415 | 194.2052 | 52.9637 | 52.9637
Staccato_22 | Staccato | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\050_Staccato_22_curves.csv | Calibration equation | 188 | 24.497 | 19.8464 | 17.329 | -28.0988 | 0.91184 | 17.1326 | 29.8589 | 0.17332 | 12.2686 | 39.8333 | 27.5646 | 27.5646 | 141.2415 | 563.0429 | 421.8014 | 421.8014
Staccato_22 | Staccato | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\050_Staccato_22_curves.csv | Offset-corrected calibration | 188 | 24.497 | 3.0598 | 2.5472 | 0.30834 | 0.85301 | -0.64561 | 6.3754 | 0.23365 | 12.2686 | 16.3013 | 4.0327 | 4.0327 | 141.2415 | 125.3727 | -15.8688 | 15.8688
Staccato_23 | Staccato | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\051_Staccato_23_curves.csv | Gradient Boosting | 173 | 22.618 | 3.1725 | 2.576 | -0.82958 | 0.89411 | 2.576 | 6.7883 | 0.14699 | 7.4519 | 11.8925 | 4.4406 | 4.4406 | 82.5737 | 141.1328 | 58.5591 | 58.5591
Staccato_23 | Staccato | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\051_Staccato_23_curves.csv | Calibration equation | 173 | 22.618 | 24.581 | 20.4462 | -108.8351 | 0.89118 | 20.368 | 38.9393 | 0.15891 | 7.4519 | 46.1864 | 38.7345 | 38.7345 | 82.5737 | 545.9723 | 463.3987 | 463.3987
Staccato_23 | Staccato | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\051_Staccato_23_curves.csv | Offset-corrected calibration | 173 | 22.618 | 7.1415 | 5.3643 | -8.2709 | 0.8169 | 3.7386 | 15.4073 | 0.25918 | 7.4519 | 22.6545 | 15.2026 | 15.2026 | 82.5737 | 167.663 | 85.0894 | 85.0894
Staccato_24 | Staccato | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\052_Staccato_24_curves.csv | Gradient Boosting | 166 | 21.78 | 4.9235 | 3.826 | 0.55829 | 0.93118 | 3.826 | 12.3539 | 0.14497 | 21.3093 | 25.7514 | 4.4421 | 4.4421 | 211.8049 | 295.5068 | 83.7019 | 83.7019
Staccato_24 | Staccato | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\052_Staccato_24_curves.csv | Calibration equation | 166 | 21.78 | 27.8213 | 23.9517 | -13.1043 | 0.90039 | 23.832 | 43.0202 | 0.18479 | 21.3093 | 60.2896 | 38.9803 | 38.9803 | 211.8049 | 733.8535 | 522.0486 | 522.0486
Staccato_24 | Staccato | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\052_Staccato_24_curves.csv | Offset-corrected calibration | 166 | 21.78 | 9.3091 | 6.8602 | -0.57911 | 0.91644 | 5.4007 | 19.4883 | 0.15621 | 21.3093 | 36.7576 | 15.4484 | 15.4484 | 211.8049 | 330.1929 | 118.388 | 118.388
Staccato_25 | Staccato | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\053_Staccato_25_curves.csv | Gradient Boosting | 166 | 21.615 | 4.2105 | 3.0213 | 0.39665 | 0.87829 | 3.0009 | 11.8261 | 0.19592 | 17.7685 | 19.4792 | 1.7106 | 1.7106 | 144.497 | 209.6708 | 65.1738 | 65.1738
Staccato_25 | Staccato | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\053_Staccato_25_curves.csv | Calibration equation | 166 | 21.615 | 8.8857 | 7.1703 | -1.6871 | 0.85235 | 6.8274 | 18.9617 | 0.21676 | 17.7685 | 27.4239 | 9.6554 | 9.6554 | 144.497 | 292.9547 | 148.4576 | 148.4576
Staccato_25 | Staccato | C:\school\project\everything everything\08b_offset_corrected_outputs\corrected_curve_data\053_Staccato_25_curves.csv | Offset-corrected calibration | 166 | 21.615 | 8.1728 | 6.3871 | -1.2732 | 0.48216 | -6.3871 | 17.473 | 0.40998 | 17.7685 | 3.892 | -13.8766 | 13.8766 | 144.497 | 5.657 | -138.84 | 138.84

## Recording-level summary
method | recordings | mean_rmse_ml_s | sd_rmse_ml_s | median_rmse_ml_s | mean_mae_ml_s | sd_mae_ml_s | median_mae_ml_s | mean_r2 | sd_r2 | median_r2 | mean_correlation_r | sd_correlation_r | median_correlation_r | mean_mean_error_ml_s | sd_mean_error_ml_s | median_mean_error_ml_s | mean_maximum_absolute_error_ml_s | sd_maximum_absolute_error_ml_s | median_maximum_absolute_error_ml_s | mean_shape_rmse_normalized | sd_shape_rmse_normalized | median_shape_rmse_normalized | mean_absolute_qmax_error_ml_s | sd_absolute_qmax_error_ml_s | median_absolute_qmax_error_ml_s | mean_absolute_volume_error_ml | sd_absolute_volume_error_ml | median_absolute_volume_error_ml
Gradient Boosting | 53 | 3.2303 | 1.4626 | 3.0457 | 2.6673 | 1.324 | 2.3877 | -1.5791 | 7.2579 | 0.45623 | 0.78166 | 0.33104 | 0.92649 | 1.9969 | 1.5096 | 1.6891 | 7.0897 | 2.6426 | 6.7883 | 0.17444 | 0.1233 | 0.1415 | 3.6886 | 2.5784 | 3.3362 | 46.2259 | 42.2302 | 37.2191
Offset-corrected calibration | 53 | 5.3989 | 2.199 | 5.246 | 4.184 | 1.7522 | 3.8884 | -1.5434 | 3.5339 | -0.46433 | 0.6385 | 0.32205 | 0.79973 | -2.3203 | 3.3639 | -2.5541 | 11.8419 | 4.8856 | 12.3874 | 0.33302 | 0.13507 | 0.29614 | 6.2917 | 4.7709 | 4.7732 | 72.1374 | 44.9309 | 72.8006
Calibration equation | 53 | 14.5069 | 6.0116 | 13.5153 | 11.9427 | 5.5572 | 10.4935 | -36.2444 | 57.4646 | -12.4961 | 0.75283 | 0.35956 | 0.90819 | 11.2061 | 5.7751 | 9.9402 | 26.6729 | 7.8676 | 26.4911 | 0.20206 | 0.13592 | 0.15891 | 23.532 | 8.3329 | 23.473 | 246.4744 | 147.812 | 200.468

## Pooled summary
method | total_time_samples | pooled_rmse_ml_s | pooled_mae_ml_s | pooled_r2 | pooled_correlation_r | pooled_mean_error_ml_s
Gradient Boosting | 8665 | 3.6461 | 2.7483 | 0.60718 | 0.87896 | 2.123
Offset-corrected calibration | 8665 | 5.6903 | 4.0577 | 0.043238 | 0.63744 | -2.1077
Calibration equation | 8665 | 15.8657 | 12.1323 | -6.4381 | 0.68973 | 11.4429